## Sheet1
Ultrasonic Unit Calibration Report---AWS
Ultrasonic Unit Model |  | USN-58L | [8 empty] | Serial Number |  |  | 01D50C
Search Unit--Size | .75" Round |  |  |  |  | Type | Straight Beam |  |  |  | Frequency |  | 2.25 |  |  |  | MHz
Calibration--Date | 2009-11-03 |  |  |  |  | Interval | 2 Months |  |  |  | Method |  | AWS D1.1
Block Serial Number |  | SN 09-3730 |  |  |  | Data | XX |  |  | As Found | [5 empty] | As Adjusted
SUPPLEMENTAL INSTRUCTIONS
Row |  |  |  | Number | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | b Average
a |  |  |  | dB Reading | 6 | 12 | 18 | 24 | 30 | 36 | 42 | 48 | 54 | 60 | 66 | 72 | 78
b |  |  |  | Display Height | 69 | 75 | 75 | 77 | 77 | 77 | 77 | 78 | 77 | 78 | 79 | 80 | 81 | 77.2857
c |  |  |  | Corrected Reading | 6.985 | 12.2608 | 18.2608 | 24.0322 | 30.0322 | 36.0322 | 42.0322 | 47.9201 | 54.0322 | 59.9201 | 65.8094 | 71.7002 | 77.5923
d |  |  |  | dB Error | -0.985 | -0.26076 | -0.26076 | -0.03217 | -0.03217 | -0.03217 | -0.03217 | 0.079908 | -0.03217 | 0.079908 | 0.19056 | 0.29982 | 0.40772
e |  |  |  | Collective db Error | 0.4 | 0.3 | 0.1 | 0.1 | 0 | 0 | 0 | -0.1 | 0 | 0.1 | 0.2 | 0.4 | 0.8
Accuracy Required: Minimum allowable range is 60dB. | bar: 78
Equipment is:
Acceptable for Use |  |  | XX |  | 12: Not Acceptable for Use; 54: Recalibration Due
Total qualified range |  |  |  | dB to | 18: dB =; 30: db; 42: Total error; 60: dB; 66: (from Chart above)
Total qualified range |  |  |  | dB to | 18: dB =; 30: db; 42: Total error; 60: dB; 66: (from form M9)
Calibrated by |  |  |  |  | 24: Level; 42: Location